## notes
you-try-it-01.xlsx
ver. | 2016-07-26
Copyright 2009-2016 Brian M. Tissue, all rights reserved.
For use with:
Brian M. Tissue, Basics of Analytical Chemistry and Chemical Equilibria, (John Wiley: New York, 2013).
http://www.achem.org
Worksheets in this file
| notes | This page with background information.
| 1.A conversions | Converting between various measurement units.
| 1.B precision | Data to find mean, standard deviation, etc.
| 1.C outlier | Test to determine if an apparent outlier may be discarded.
| 1.D LOD-LOQ | Determining limit of detection and limit of quantitation.
| 1.E calibration | Calculations based on one-point calibration.
| 1.F standard-addition | Calibration using standard addition.
Background
Refer to Chapter 1 in the text for equations and explanations.
Each worksheet has instructions in the blue shaded box.
For step-by-step help see you-try-it-01guide.pdf.
For help getting started with Excel see spreadsheet-help.pdf.

## 1.A conversions
You-Try-It 1.A |  |  | Unit Conversions
1. | Table 1.A.1 provides several examples of physical measurements in common units.
Express each case in Table 1.A.1 in the other units that are listed.
Some conversion factors are below the table.
2. | Table 1.A.2 provides a list of analytical results with common units.
Express each case in Table 1.A.2 in the other units that are listed.
The density of each solution is given in column E.
Notes: | mph is 'miles per hour'
‰ is 'parts per thousand' (by weight if not specified otherwise)
Table 1.A.1
sample |  |  | measurement | value | units | Convert to:
vehicle |  |  | velocity | 25 | mph |  | km/h |  | m/s
10 % sucrose solution |  |  | density | 1.034 | g/cm3 |  | g/mL |  | kg/m3
7Ag85Cu8Sn brazing alloy |  |  | density | 4.8 | oz/cu in |  | g/mL |  | kg/m3
|  | value | units | value | units | value | units | value | units
|  | 1.6093 | km/mi | 1000 | g/kg | 1 | mL/cm3 | 10000 | ppm/%
|  | 28.3495 | g/oz | 100 | cm/m | 1000 | mL/L | 1000 | ppm/‰
|  | 16.3871 | mL/cu in | 1000 | mm/m | 1000 | L/m3
Table 1.A.2
analysis |  |  | density (g/mL) | value | units | Convert to:
nitrate in water |  |  | 1 | 1.9 | ppm |  | wt % |  | M
NaCl in seawater |  |  | 1.025 | 30 | ‰ |  | ppm |  | M
acetic acid in vinegar |  |  | 1.01 | 5.5 | % by wt |  | ppm |  | M
ethanol in beverage |  |  | 1.02 | 4.2 | % (v/v) |  | wt % |  | M
|  |  | species | formula wt |  | density
|  |  | NO3− | 62.0049 | g/mol
|  |  | NaCl | 58.4428 | g/mol
|  |  | CH3COOH | 60.052 | g/mol
|  |  | C2H5OH | 46.0684 | g/mol | 0.789 | g/mL

## 1.B precision
You-Try-It 1.B |  | Precision |  |  |  |  | Table of t values for given alpha levels
[6 empty] | 0.2 | 0.1 | 0.05 | 0.025 | 0.01 | 0.001
1. | Calculate the wt-% for each measurement in Table 1.B.1. |  |  |  | N-1 | (80%) | (90%) | (95%) | (97.5%) | (99%) | (99.9%)
| Find the mean, standard deviation, and %-RSD. |  |  |  | 1 | 3.078 | 6.314 | 12.706 | 31.821 | 63.657 | 636.6
[5 empty] | 2 | 1.886 | 2.92 | 4.303 | 6.965 | 9.925 | 31.59
2. | Column E contains formulas for statistical results for up to 20 data points in Table 1.B.2. |  |  |  | 3 | 1.638 | 2.353 | 3.182 | 4.541 | 5.841 | 12.92
| Write formulas to calculate the variance and standard error. |  |  |  | 4 | 1.533 | 2.132 | 2.776 | 3.747 | 4.604 | 8.61
| Use the table of t values to the right to find the 90 and 99 % confidence levels. |  |  |  | 5 | 1.476 | 2.015 | 2.571 | 3.365 | 4.032 | 6.869
| Verify your results using Data -> Data Analysis -> Descriptive Statistics. |  |  |  | 6 | 1.44 | 1.943 | 2.447 | 3.143 | 3.707 | 5.959
[5 empty] | 7 | 1.415 | 1.895 | 2.365 | 2.998 | 3.5 | 5.408
[5 empty] | 8 | 1.397 | 1.86 | 2.306 | 2.896 | 3.355 | 5.041
[5 empty] | 9 | 1.383 | 1.833 | 2.262 | 2.821 | 3.25 | 4.781
[5 empty] | 10 | 1.372 | 1.812 | 2.228 | 2.764 | 3.169 | 4.587
Table 1.B.1 Protein analysis of extra crunchy peanut butter. |  |  |  |  | 15 | 1.341 | 1.753 | 2.131 | 2.602 | 2.947 | 4.073
[5 empty] | 20 | 1.325 | 1.725 | 2.086 | 2.528 | 2.845 | 3.85
Trial | Sample wt (g) | Protein wt (g) | wt-% |  | 25 | 1.316 | 1.708 | 2.068 | 2.485 | 2.787 | 3.725
1 | 9.91 | 1.92 | 19.3744 |  | 30 | 1.31 | 1.697 | 2.068 | 2.457 | 2.75 | 3.646
2 | 10.17 | 2.04 | 20.059 |  | 50 | 1.299 | 1.676 | 2.068 | 2.403 | 2.678 | 3.496
3 | 10.54 | 2.09 | 19.8292 |  | 100 | 1.29 | 1.66 | 2.068 | 2.364 | 2.626 | 3.391
4 | 10.01 | 2.05 | 20.4795 |  | infinity | 1.31 | 1.645 | 2.068 | 2.326 | 2.576 | 3.3
|  | mean: |  | %
|  | std dev: |  | %
|  | relative standard deviation: |  | %-RSD
Table 1.B.2 Descriptive Statistics
Trial | Data |  | Statistics
1 | 3.44 | N = | 5
2 | 3.11 | sum = | 15.83
3 | 2.98 | mean = | 3.166
4 | 3.27 | std dev = | 0.18849
5 | 3.03
6 |  | %-RSD =
7 |  | variance =
8 |  | std error =
9 |  |  |  | t values
10 |  | 90-% C.L. =
11 |  | 95-% C.L. = | 0.23401 | 2.776
12 |  | 99-% C.L. =
13
14
15
16
17
18
19
20

## 1.C outlier
You-Try-It 1.C |  | Discarding an Outlier |  |  |  |  | Critical values of Dixon's Q parameter (Qc)
[7 empty] | N | 95% | 99%
1. | Determine which data value in Table 1.C.1 is a potential outlier and calculate Q. | [5 empty] | 3 | 0.97 | 0.994
| Determine if the outlier may be discarded at the 95-% or 99-% confidence levels. | [5 empty] | 4 | 0.829 | 0.926
[7 empty] | 5 | 0.71 | 0.821
2. | Repeat using Peirce's criterion. | [5 empty] | 6 | 0.625 | 0.74
[7 empty] | 7 | 0.568 | 0.68
[7 empty] | 8 | 0.526 | 0.634
[7 empty] | 9 | 0.493 | 0.598
[7 empty] | 10 | 0.466 | 0.568
Table 1.C.1 Calcium Potentiometry | [6 empty] | 15 | 0.384 | 0.475
Trial | ISE Result (mV) |  | Results |  |  |  | 20 | 0.342 | 0.425
1 | 39.8 | N = | 5 | Qc (95 %) = | 0.71 |  | 25 | 0.317 | 0.393
2 | 36.5 | sum = | 195 | Qc (99 %) = | 0.821 |  | 30 | 0.298 | 0.372
3 | 39.9 | mean = | 39
4 | 39.2 | std dev = | 1.423 |  |  |  | suspect may be rejected if Q > Qc
5 | 39.6
A more complete list of Qc values is in:
David B. Rorabacher,
Dixon Q-test calculation | [6 empty] | "Statistical treatment for rejection of deviant values: critical values of Dixon's "Q"
| value | deviation | closest value | Dixon's Q | Qc (95 %) | Qc (99 %) | parameter and related subrange ratios at the 95% confidence level,"
min = | [6 empty] | Anal. Chem., 1991, 63(2), 139-146; DOI: 10.1021/ac00002a010.
max =
Values of R for Peirce's Criterion
Peirce's criterion | [7 empty] | Number of doubtful observations
| value | ¦deviation¦ | R | dmax | result |  | N | 1 | 2 | 3 | 4
min = | [6 empty] | 3 | 1.196
max = | [6 empty] | 4 | 1.383 | 1.078
[7 empty] | 5 | 1.509 | 1.2
[7 empty] | 6 | 1.61 | 1.299 | 1.099
[7 empty] | 7 | 1.693 | 1.382 | 1.187 | 1.022
[7 empty] | 8 | 1.763 | 1.453 | 1.261 | 1.109
[7 empty] | 9 | 1.824 | 1.515 | 1.324 | 1.178
[7 empty] | 10 | 1.878 | 1.57 | 1.38 | 1.237
[7 empty] | 11 | 1.925 | 1.619 | 1.43 | 1.289
[7 empty] | 12 | 1.969 | 1.663 | 1.475 | 1.336
[7 empty] | 13 | 2.007 | 1.704 | 1.516 | 1.379
[7 empty] | 14 | 2.043 | 1.741 | 1.554 | 1.417
[7 empty] | 15 | 2.076 | 1.775 | 1.589 | 1.453
[7 empty] | 20 | 2.209 | 1.914 | 1.732 | 1.599
[7 empty] | 25 | 2.307 | 2.019 | 1.84 | 1.709
suspect may be rejected if ¦deviation¦ > dmax
A more complete list of R values is in:
Stephen M. Ross,
"Peirce's criterion for the elimination of suspect experimental data,"
Journal of Engineering Technology, Fall 2003.

## 1.D LOD-LOQ
You-Try-It 1.D |  | Limit of Detection (LOD) and Limit of Quantitation (LOQ)
Tables 1.D.1 and 1.D.2 contain data sets from two different fluorescence experiments.
The fluorescence signal is directly proportional to analyte concentration for both data sets.
The blank measurements are for pure solvent.
1. | Determine the S/N, LOD, and LOQ for the measurement in Table 1.D.1.
2. | Correct the measurements in Table 1.D.2 for the nonzero background and plot.
Determine the LOD and LOQ.
Table 1.D.1. Repetitive fluorescence measurements.1
| blank | standard
trial | (cps) | (cps) |  | standard concentration: | 10 | ppm
1 | 753 | 1667287 |  |  | 1000 | ppb/ppm
2 | 620 | 1670694 |  | 3x noise = |  | cps
3 | 703 | 1670408 |  | 10x noise = |  | cps
4 | 620 | 1668215 |  | calibration = |  | cps/ppb
5 | 576 | 1927720
average |  |  |  | LOD: |  | ppb
std dev |  |  |  | LOQ: |  | ppb
Table 1.D.2. Fluorescence signal as a function of concentration.2
conc | corrected | measured | blank
(mM) | signal (mV) | signal (mV) | signal (mV)
0.3 | 0.581 | 0.581 | 0.006
0.094868 | 0.173 | 0.173 | 0.003
0.03 | 0.059 | 0.059 | 0.006 | 3x noise =
0.0094868 | 0.029 | 0.029 | 0.005 | 10x noise =
0.003 | 0.01 | 0.01 | 0.006 | slope =
0 | 0.007 | 0.007 | 0.004
|  | average |  | LOD: |  | mM
|  | std dev |  | LOQ: |  | mM
1. The fluorescence signal is counts per second (cps) from a photomultiplier tube detector and pulse counter.
2. The signal is a voltage from a photon detector and lock-in amplifier.

## 1.E calibration
You-Try-It 1.E |  | Calibration |  |  | Measurement Procedure
Crush an analgesic tablet and weigh ≈ 0.1 g to four places on an analytical balance.
Acetylsalycilic acid (aspirin or ASA) can be measured using UV light absorption (see procedure). |  |  |  |  | Dissolve in deionized water and allow time for all of the ASA to dissolve (do not heat).
The measured absorbance is directly proportional to ASA concentration: A µ cASA. |  |  |  |  | Filter to remove the starch binder and wash with water.
Table 1.E.1 lists absorbance measurements for a standard solution and two unknowns. |  |  |  |  | Dilute to a known volume with 0.05 M HCl.
Measure the absorbance of the solution at 227 nm.
1. | Calculate the ASA concentration in the unknown samples using a simple proportionality.
A standard ASA solution of 5.0x10−5 M in 0.05 M HCl has an absorbance of 0.363 at 227 nm.
2. | Generate a two-point calibration curve using 0,0 and the standard measurement.
| Calculate the ASA calculation in the unknown samples using the calibration curve. |  |  |  | Sample 2 exhibited a faint turbidity (cloudiness). It was remeasured at a wavelength where
ASA does not absorb to provide an approximate absorbance due to the turbidity
This measurement is listed as the blank measurement for Sample 2.
Table 1.E.1
|  | absorbance | blank | conc (M)
| standard | 0.363 | 0 | 5e-05
| Sample 1 | 0.222 | 0
| Sample 2 | 0.311 | 0.043
|  | conc (M) | absorbance
| blank | 0 | 0
| Standard | 5e-05 | 0.363

## 1.F standard-addition
You-Try-It 1.F |  | Standard Addition
Table 1.F.1 gives a set of data for lead analysis using an electrochemical method.
The dependence is linear, i.e., electrochemical current is proportional to analyte concentration.
1. | Determine the unknown concentration using a proportionality calculation.
2. | Determine the x-intercept for the data set, which equals the unknown concentration.
Table 1.F.1
| std addition | signal | cunk
| (μM) | (μA) | (μM)
| 0 | 2.4
| 2.5 | 5.2
| 5 | 8.2
| 7.5 | 11
|  | average: |  | µM
std dev:
data adapted from:
Andrew J. Saterlay , Shelley J. Wilkins and Richard G. Compton
"Towards greener disposal of waste cathode ray tubes via ultrasonically enhanced lead leaching"
Green Chem., 2001, 3, 149 - 155, DOI: 10.1039/b102671m
http://www.rsc.org/ej/GC/2001/b102671m/